[image: image1.jpg]a "Associated
o B Power Technologies



[image: image1.jpg]





QUICK START GUIDE 300XAC SERIES
1. Input Power

On the rear panel of the instrument wire the appropriate line voltage to the input terminal block.  Refer to the manual for the correct voltage, current and wire gauge per the specific 300XAC Series model that you are using.

2. Attach Linking Cables (Only if linking sources)

On the rear panel of the instrument if the linking card option has been installed interconnect the linking cables from the linking output to the linking input connectors.  If you are connecting two sources connect the linking cable from the linking output of one source to the linking input of the second source.  Only one linking cable is used with two sources.  

If you are linking three sources connect the linking cable from the linking output of one source to the linking input of the second source.  On the second source connect the linking cable from the linking output to the linking input of the third source.  Two linking cables will be used in this configuration and the second source will have a cable attached to the linking output and linking input.
3. Powering On

Power on the instrument with the power switch located on the left hand side of the instrument.

4. Selection of MANUAL Mode

Press the <more> soft key until the soft key System is shown.  Press the System soft key and use the up or down arrow soft key to highlight the parameter Auto Run.  Now press the Edit soft key and use the Change soft key to select the Auto Run mode to MANUAL.  Now press the Enter soft key and then the Exit soft key two times.
5. Selection of FUNCTION Mode (Only if linking sources)
Press the <more> soft key until the soft key System is shown.  Press the System soft key and use the up or down arrow soft key to highlight the parameter FUNCTION.  Now press the Edit soft key and use the Change soft key to select the FUNCTION mode to Parallel, 1Φ3W, or 3Φ4W.  Now press the Enter soft key and then the Exit soft key two times.  The FUNCTION mode changes the output voltage configuration.

6. Editing Memory, Voltage, Frequency, & Current High Limit
To edit memory, voltage, frequency, & current high limit press the Edit soft key.  Use the up or down arrow soft key to highlight the parameter you would like to edit and press the Edit soft key.
A. Editing Memory
Press the <more> soft key to get to the Name soft key and press the Name soft key.  This will provide the character map to enter the memory name.  Press the <top> soft key to use the soft keys >, down arrow, and Select to choose your characters.  The numeric keypad is also available for character entry.  Use the <more> soft key to provide the Enter soft key for editing the memory name.  Press the Esc soft key one time to get back to selecting the memory for editing.  Press the Exit soft key a two times to leave the editing screen.

B. Editing Voltage

The voltage parameter will be highlighted.  Use the numeric keypad to enter the voltage and press the Enter soft key to accept the value.  Press the Exit soft key one time to get back to selecting the parameter for editing. Press the Exit soft key a second time to leave the editing screen. Within the voltage parameter you can also press the Change soft key to toggle the voltage mode (HIGH/AUTO).  Press the Enter soft key to accept the voltage output mode.

Voltage can also be edited instantaneously via the rotary knob on the front panel when the output has been initialized.  Spin the rotary knob clockwise to increase the voltage and counterclockwise to decrease the voltage. 

C. Editing Frequency

The frequency parameter will be highlighted.  Use the numeric keypad to enter the frequency and press the Enter soft key to accept the value.  Press the Exit soft key one time to get back to selecting the parameter for editing. Press the Exit soft key a second time to leave the editing screen.

Frequency can also be edited instantaneously via the rotary knob on the front panel when the output has been initialized.  Spin the rotary knob clockwise to increase the frequency and counterclockwise to decrease the frequency.  The meter selection must be on F in order to change the frequency in this mode.  To change the meter display press the Meter soft key. 

D. Editing Current High Limit
The current limit parameter will be highlighted.  Use the numeric keypad to enter the current high limit.  Press the Exit soft key one time to get back to selecting the parameter for editing. Press the Exit soft key a second time to leave the editing screen.

7. Initializing/Terminating Output
To initialize or terminate the output press the TEST/RESET key.  When the LED is in a solid state the output is active.

8. Changing Display Metering

When the output is initialized the Meter soft key becomes available, which allows you to change the metering for the right side of the display.  Press the Meter soft key will toggle the metering from frequency, current, power, reactive power, crest factor, peak current, apparent power, and power factor.
F:\APT\Training&Procedures\300XAC\QuickStartGuide300XACSeries.doc













PAGE  
1
Associated Power Technologies, Inc.


1142 S. Diamond Bar Blvd # 106, Diamond Bar, CA 91765 | Toll Free: +1-877-322-7693 | Phone: +1-909-860-1646
Fax: +1-909-860-2727 | Email: info@aspowertechnologies.com | Web: www.aspowertechnologies.com

