## Slide 1
• Ground of electrical products
• Class I vs. Class II products
• Ground Continuity Test
• Ground Bond Test
• What is tested during each test
Testing the Ground Circuit

## Slide 2
Meet Our Team

## Slide 3
Webinar Notes
Please use the Q & A utility to ask us any questions concerning the material being presented.
You can find a recording of this webinar and presentation on our Webinar Archive page, https://www.arisafety.com/webinars/archived-webinars/
Please contact Brittany Socha – on the chat line or email Brittany.socha@ikonixusa.com if you have any connection issues.

## Slide 4: Ground Bond & Ground Continuity Testing Learning Objectives
Ground
Ground Bond
Ground Continuity
Bond vs. Continuity
What is Ground?
Grounding of electrical products
Class I vs. Class II products
The Ground Bond Test – What is it?
Ground Bond Standards Information
Ground Bond Testing Examples
The Ground Continuity Test – What is it?
Ground Continuity Standards Information
Ground Continuity Testing Examples
Comparing the Ground Bond and Ground Continuity tests
Example of Ground Bond vs Ground Continuity
Notes: Important to understand what a ground

## Slide 5
Ground or “earth” can have numerous functions and meanings. The main purpose of a ground is to reduce the risk of electrical shock.
1 – http://en.wikipedia.org/wiki/Ground_%28electricity%29#mediaviewer/File:HomeEarthRodAustralia1.jpg
2 - http://www.howequipmentworks.com/physics/electricity/elec_safety/electrical_safety.html
Reference point for an electrical circuit
Second line of defense against electrical shock1
Common return path for electrical current2
What is Ground?

## Slide 6
Modern Power Distribution

## Slide 7
System or Service ground – designed to protect machines, tools and insulation.
Equipment Ground – Designed to protect operator from electrical shock hazards.
Service Ground
Equipment Ground
Ground or “Earth”
Different Functions of Ground

## Slide 8
Class I products provide electrical protection by employing basic insulation and a grounded chassis. Class II products employ a double layer of insulation so a grounded chassis is not necessary.
Class I vs. Class II Electrical Products

## Slide 9
What is Good Ground?
NFPA (National Fire Protection Agency) stipulates the NEC which is adopted in all 50 U.S. States. The NEC gives requirements for grounding products and installations. The NEC defines a “good” ground.
National Electrical Code
NEC 250-45
NEC 250-51
NFPA 70
Protect people and property from electrical hazards
”Any exposed non-current carrying metal parts of cord & plug connected equipment which may become energized shall be grounded.”
Permanent and Continuous
Capacity to conduct fault current
Low impedance to limit voltage to ground.

## Slide 10
Ground Bond and Continuity Tests
• Ground Bond are generally TYPE TESTS.
• Ground Continuity are generally PRODUCTION TESTS.
These two types of tests are used to prove and verify that there is a low impedance path to the ground for electrical products and systems. This ensures proper grounding.
Purpose
Method
Testing
Ground Bond and Ground Continuity tests are designed to “check” the ground connection on a product or system.
How you ensure you have a proper ground.
Apply a current to the ground point of a product or system.
Measure potential drop across the ground circuit and calculate the circuit resistance.
Permanent and Continuous
Capacity to conduct fault current
Low impedance to limit voltage to ground.

## Slide 11
Quiz Question
Which of the following best represents the result of a ground Bond test?

## Slide 12
Associated Research HYAMP connected to
a DUT (Device Under Test)
The Ground Bond Test

## Slide 13
Ground Bond Testing

## Slide 14
Ground Bond Testing

## Slide 15
Ground Bond Testing

## Slide 16
By the Numbers – What the Standards State
IEC/UL 60601-1 3rd Edition

## Slide 17
By the Numbers – What the Standards State
UL 1598/CSA C22.2 No. 250.0-08 3rd Edition Luminaires

## Slide 18
IEC/UL 60335-1 5th Edition
By the Numbers – What the Standards State

## Slide 19
Ground Continuity Testing
Ground Continuity testing is the same concept as Ground Bond testing. You’re simply checking for a continuous path on a ground circuit. However, there are key differences between these two tests.
Purpose
Method
A DC low current test to check circuit resistance.
Similar concept to the Ground Bond testing.
Not all readings display on Ohmic value
Light/buzzer meters to check whether continuity exists.

## Slide 20
IEC/UL 61010-1 3rd Edition
By the Numbers – What the Standards State

## Slide 21
IEC/UL 60335-1 5th Edition
By the Numbers – What the Standards State

## Slide 22: Ground Bond Vs. Ground Continuity
The Ground Bond test is more stringent than the Ground Continuity test.
Ground Continuity Test | Ground Bond Test
• Verifies the existence of a ground connection • Readings generally given in Ωs • The test is quick to set up and easy to perform • Usually used as an extra feature during the Hipot test. | • Verifies the integrity of a ground connection • Readings generally given in mΩ • Provides more valuable safety information about DUT • Can be combined with a Hipot test for a more complete safety testing system.

## Slide 23
Example: 64 strand wire with all but one strand connected.
Ground Bond Vs. Ground Continuity

## Slide 24
Example: 64 strand wire with all but one strand connected. A ground continuity test would pass this wire because the instrument reads a continuous path.
Ground Bond Vs. Ground Continuity

## Slide 25
Example: A Ground Bond test load the wire with high current. This would “burn up” the connection and cause a FAIL.
Ground Bond Vs. Ground Continuity

## Slide 26
Poll Question
When are you performing the Ground Bond test on your product?

## Slide 27
Video Demonstration

## Slide 28
Visit us online to view all of our Educational Resources
arisafety.com/support/educational-resources/
Educational Resources

## Slide 29
An Introduction to Power Sources (APT)
Our Next Webinar is
Join Us
Wednesday, September 12 at
10AM CST (GMT – 6:00) Click the link in chat box to learn more
Notes: Insert next webinar info here

## Slide 31
You can find a recording of this webinar and presentation on our Webinar Archive page, https://www.arisafety.com/webinars/archived-webinars/
Check out our website for more information arisafety.com
For any questions about our webinars send an email to Brittany at brittany.socha@ikonixusa.com
Contact Us